## AI validation
Ticket | Column2 | Symptom | Root Cause | Person in charge | Symptom/Root Cause correct | If AI Jira-Bug to Symptom assignment is not correct, which symptom and Root Cause fits better | Bug Cluster | Comment | Assigne | Retesting necessary | Fix can be implemented. | Fix can be implemented in which duration (weeks) | Necessary changes for this fix (new IOSS, new FW, new FPGA Version, new Hardware,..) | Estimated impact (Level 1:low-5:high)
Y3839-231 | Schick AE sensor connection LED sporadically not powered when the USB 3.0 connection to host PC is interrupted | Intermittent Connectivity | Reconnection of sensor not stable | Christian | yes |  |  |  | Xiang Ding | yes
Y3839-234 | Connection Status LED lights up dimly when USB3 remote plugged in host PC | No Power / "Dying Boxes" (USB module) | USB3 remote hardware failure | Guido | yes |  |  | Reassign to FPGA | Xiang Ding
Y3839-255 | Missing message when starting intraoral acquisition with insufficient disk space | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | Yes |  |  | Team_DX5
Y3839-265 | USB3 Remote LED interfaces have 3 color states | Intermittent Connectivity | Loose USB3 remote/sensor interface | Christian | no | cannot be categorized, LED sequence is due to the boot process |  |  | Sebastian Torpus | no | no, due to boot process | n.a. | n.a. | 1
Y3839-282 | 8000991426, XIOS AE sporadically no image provided | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | No fastboot support in IOSS |  | Device Resume Fast Boot | Jannik Kappes | no | yes | 1 week (cluster) | new IOSS | 4
Y3839-304 | Sporadically no sensor initialization phase at Exposure view | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue | Erik | Yes |  |  |  | Team_DX5 | no | yes | 0.3 weeks | new IOSS | 3
Y3839-305 | Acquisition view shows wrong loading indicator during exposure with sensor disconnect | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Team_DX5
Y3839-307 | Sensor upgrading fails for XIOS XG | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | Sensor upgrading not implemented |  |  | Sebastian Torpus | no | no, only one sensor FW exists | n.a. | n.a. | 1
Y3839-308 | Missing holder position aid pictures on Endo Horizontal (15) | Unassigned |  | Christian | not assigned | graphic (positioning holder) missing in Sidexis |  |  | Team_DX5
Y3839-309 | Unsupported language template list output inconsistently translated | Unassigned |  | Christian | not assigned | translation issue in Sidexis |  |  | Team_DX5
Y3839-340 | Documentation: Missing document proof for test Sensor IEC and EN Performance Standards | Unassigned |  | Christian | not assigned | documentation |  |  | Christian Beckhaus |  | yes | 2 | documentation only
Y3839-373 | Disconnecting sensor from USB2-Remote does not reliably indicate the missing sensor in the status bar | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Team_DX5
Y3839-382 | Sporadic "Application Error" when starting exposure series | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync) | Erik | No | IOSS bugs |  | possible race condition | Team_DX5
Y3839-383 | Connected device is not recognized by IOSS | Intermittent Connectivity | Loose micro-B USB interface (no cable clip used) | Christian | yes |  | Device connection race condition |  | Team_DX5 | no | yes | 0.7 weeks | new IOSS | 4
Y3839-384 | Insufficient buff size resulting in error during exposure and loss of image | Images not transferred | SW allocates too small buffer size | Steffen | Yes | IOSS bugs |  |  | Team_DX5
Y3839-422 | Applicable standards missing | Interface Update Issues | Others currently unknown | Michael | no | cannot be categorized, Docu bug |  |  | Frank Darmer
Y3839-425 | Outdated standards in test | Unassigned |  | Christian | not assigned | standards in test |  |  | Christian Beckhaus |  | yes | 2 | documentation only
Y3839-426 | Missing evidence for UL 60601-1 and CAN/CSA C22.2 No.601.1-M90 | Unassigned |  | Christian | not assigned | CB report |  |  | Christian Beckhaus |  | yes | 2 | documentation only
Y3839-471 | Remote signals sensor connection even tokens are not received any more | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Sebastian Torpus | no | yes | quick but needs broad testing | new FW
Y3839-478 | Acquisition app not initializing properly | Intermittent Connectivity | IOSS bugs | Christian | yes |  |  |  | Team_DX5 | no | yes | 0.3 weeks | new IOSS | 3
Y3839-479 | Acquisition App not able to fetch sensor information and device becomes unavailable | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | ?
Y3839-506 | IOSS enables Sensor Firmware update even if the firmware is up to date | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | IOSS bugs |  |  | Team_DX5
Y3839-510 | Only one firmware available for XIOS XG Supreme/Schick 33 sensors | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | Only one sensor firmware |  |  | Christian Beckhaus | no | no | N/A | new sensor firmware (out of scope)
Y3839-511 | X-ray tab in Intraoral Config app only support German and English | Unassigned |  | Christian | not assigned | translation issue in Intraoral Config App |  |  | Team_DX5
Y3839-513 | Performing an exposure causes an unhandled state when connecting IOSS to a Server different from the one the local Sidex | Unassigned |  | Christian | Intermittent Connectivity | Wrong IOSS configuration |  | Proposal from Steffen | Team_DX5
Y3839-515 | Documentation: Binning is not taken into account in requirement for visualization time | Unassigned |  | Christian | not assigned | documentation |  |  | Team_DX5
Y3839-520 | Saving image did not succeed | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | IOSS bugs | Memory Leak Save Error |  | Team_DX5 | no | yes | 1 week (cluster) | new IOSS | 5
Y3839-521 | USB Device not detected reliably after PBT_APMRESUMEAUTOMATIC | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  | Device Resume Fast Boot |  | Team_DX5 | no | yes | 1 week (cluster) | new IOSS | 3
Y3839-522 | IOSS Crashes when disconnecting supreme sensor from USB 2 interface | Intermittent Connectivity | IOSS bugs | Christian | yes |  |  |  | Team_DX5
Y3839-523 | In the acquisition window acquired Image would not be displayed if during image acquisition the selected image odontogra | Previous (Patient) Image | Missing error handling in SW/FW | Steffen | Yes | IOSS bug |  |  | Team_DX5
Y3839-524 | During Bensheim automated testing - randomly Autotake flashes red even though everything is fine and an image is still a | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync) | Erik | Yes |  | State Machine Consistancy FW/SW |  | Team_DX5 | no | yes | 3 weeks (cluster) | new IOSS, new FW | 5
Y3839-525 | Remote and sensor stuck in ready, not switching to ready-for-exposure | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue | Erik | Yes |  |  |  | Sebastian Torpus | yes (with device logs)
Y3839-526 | Re-Plug-in of sensor not handled correctly | Intermittent Connectivity | Reconnection of sensor not stable | Christian |  | --> in my opinion solved ? |  |  | Sebastian Torpus | yes
Y3839-529 | Saving image failed due to resources ; Image Loss | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | IOSS bugs | Memory Leak Save Error |  | Team_DX5 | no | yes | 1 week (cluster) | new IOSS | 5
Y3839-530 | RED LED State is not defined in IFU and Designs | Intermittent Connectivity | Missing error handling in SW/FW | Christian | no | cannot be categorized, IFU/docu topic |  |  | Frank Darmer
Y3839-533 | Sensor never becomes exposure ready when calibration file read fails | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue | Erik | No | FW / SW Error handling |  | calib file corrupted | Team_DX5
Y3839-537 | InUse lock not released when SSE connection between Sidexis and IO-Sensor-Service is lost | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Team_DX5 | no | yes | 0.3 weeks | new IOSS | 3
Y3839-538 | USB race condition on cold boot causes sensor NoHardware state | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  | Device connection race condition |  | Team_DX5 | no | yes | 0.7 weeks | new IOSS | 4
Y3839-541 | Negative effect of checkerboard pattern on image quality | Overexposed images | Error in the image processing path | Xiang | No |  |  |  | Xiang Ding
Y3839-545 | USB3 remote was stuck in boot loop | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue | Erik | Yes | State-machine mismatch in SW/FW (out-of-sync) |  |  | Team_DX5
Y3839-546 | Exposure appears double in template acquisition | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue | Erik | Yes | State-machine mismatch in SW/FW (out-of-sync) | State Machine Consistancy FW/SW |  | Steffen Wenz | no | yes | 3 weeks (cluster) | new IOSS, new FW | 5
Y3839-547 | Spring Oaks -> Lost Images | Images not transferred | Software timing error | Steffen | Yes | Yes | State Machine Consistancy FW/SW |  | Steffen Wenz | no | yes | 3 weeks (cluster) | new IOSS, new FW | 5
Y3839-549 | defective sensors @practices Rümenapp and Ernst | Loosening Screws | Incorrect torque used during sensor cable exchange | Guido | yes |  |  |  | Christian Kurz | yes | no | - | maybe HW topic or a general Gen 1 topic -> has to be checked in more detail by C. Kurz and G. Scheibl
Y3839-555 | X-ray applied but no image acquired/transmitted (Spring Oaks) | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | Sensor self-triggering | State Machine Consistancy FW/SW | Symptom might not be correct, but it is a case of self-triggering | Team_DX5 | no | yes | 3 weeks (cluster) | new IOSS, new FW | 5
Y3839-567 | Eaglesoft/IOSS: USB3 interface not ready for next exposure though Eaglesoft indicates so | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue | Erik | Yes | State-machine mismatch in SW/FW (out-of-sync) |  | Duplicates Y3839-570 | Christian Kurz (changed to Team_DX5) | no | yes, but discussed with SW because issue is strongly related to a new IOSS version (arming of sensor has to be removed after a X-Ray shot -> a new SetEventDetection has always be send to the interface/state 4 always necessary (checkered pattern is written in the HyperRAM) -> correct signaling is reached -> X-Ray is triggered => breaking changes, backward compatibility is not given any more | ? | new FW, new IOSS and new third party integration | 5
Y3839-568 | Sidexis/IOSS: Viewport remains red after sensor reconnect, sensor shown twice in configuration app and Sidexis | Intermittent Connectivity | Reconnection of sensor not stable | Christian | no | Intermittent Connectivity -> Missing Error handling in SW/FW | Device connection race condition |  | Team_DX5
Y3839-570 | Eaglesoft/IOSS: No image received after exposure, USB3 interface stuck in state 6 | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | State-machine mismatch in SW/FW (out-of-sync) |  | handling in IOSS possible ->DX5? | Christian Kurz (changed to Team_DX5) | no, similar to 567 | yes, but discussed with SW because issue is strongly related to a new IOSS version (arming of sensor has to be removed after a X-Ray shot -> a new SetEventDetection has always be send to the interface/state 4 always necessary (checkered pattern is written in the HyperRAM) -> correct signaling is reached -> X-Ray is triggered => breaking changes, backward compatibility is not given any more | 3 weeks (cluster) | new FW, new IOSS and new third party integration | 5
Y3839-571 | Eaglesoft/Dentrix [CDR]: "Unable to take X-Ray" error after switching to AE sensor, no image acquired | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | Yes | State-machine mismatch in SW/FW (out-of-sync) |  |  | Christian Kurz
Y3839-573 | Dentrix/CDR: USB3 interface fails to reconnect to sensor after multiple plug cycles, sensor LED stays orange | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | Yes |  |  | CDR not IOSS | Rami Alshamma | no | no |  | is a CDR specific issue
Y3839-575 | Sensor sporadically not detected after repeated USB3 Interface reconnect | Intermittent Connectivity | Reconnection of sensor not stable | Christian | yes |  |  |  | Christian Kurz | yes (with device logs)
Y3839-583 | Dentrix/CDR: No image delivered after exposure following ~9 min wait in ready state | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | Yes |  |  | CDR not IOSS | Christian Kurz | no | no |  | is a CDR specific issue
Y3839-586 | White images with different sensors (Waukee Dental) | White Images | Sensor self-triggering | Xiang | Yes |  |  |  | Sebastian Torpus | no | yes, but discussed with SW (only for exposure readiness) | 1 week | new FW, potentially new IOSS
Y3839-587 | SW freezes after taking some images (Waukee Dental) | Intermittent Connectivity | Missing error handling in SW/FW | Christian | no | cannot be categorized, multiple issues |  |  | Sebastian Torpus
Y3839-588 | Multiple different issues with one sensor (Waukee Dental) | Intermittent Connectivity | Defect USB3 remote (cold solder joints) | Christian | yes |  |  |  | Sebastian Torpus
Y3839-593 | Checkered pattern image in a X-Ray series (Dr. Mathes) | Images not transferred | Missing error handling in SW/FW | Steffen | Yes | Read error 8 |  |  | Xiang Ding
Y3839-594 | Dentrix/CDR: The Sensor LED stays orange sporadically after unplugging the sensor | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Xiang Ding
Y3839-596 | Sensor triggers due to electrostatic charge without x-ray and shows white image | White Images | Sensor triggers by the influence of an electric field | Xiang | Yes |  |  |  | Markus Damnik
Y3839-598 | USB3 Interface Stuck in Red LED State which requires PC Restart to Restore Connection | Intermittent Connectivity | Driver conflicts (CDR vs IOSS) | Christian | yes |  |  |  | Sebastian Torpus | seemingly very rare to occur | no
Y3839-599 | Sidexis[IOSS]/Dentrix[CDR]: USB3 Interface Sporadically Fails to Detect Sensor After Reconnection | Intermittent Connectivity | Reconnection of sensor not stable | Christian | yes |  |  |  | Xiang Ding
Y3839-607 | USB3 Interface: Sensor LED turns orange during consecutive acquisitions in TA | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Michael Petek | yes (with device logs)
Y3839-611 | FirmwareUpgradeUtility is sending commands when an USB Stick is connected | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | Firmware Upgrader bug |  |  | Steffen Wenz
Y3839-625 | GETFPGAVersions returns null after successful flash | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | FPGA flash erase issue |  |  | Xiang Ding
Y3839-631 | Unexpected host LED state in non-Sidexis setup | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | No | FW Led state bug |  |  | Team_DX5
Y3839-638 | Dentrix/CDR: Dentrix shows "Ready for exposure" while Interface LED is orange | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | Yes | State-machine mismatch in SW/FW (out-of-sync) |  |  | Team_DX5
Y3839-639 | It is possible to acquire images with IOSS without enabling event detection | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | No | Exposure ready FW / SW bug | State Machine Consistancy FW/SW | enables misuse | Steffen Wenz | no | yes | 3 weeks (cluster) | new IOSS, new FW | 5
Y3839-657 | Dentrix [CDR]: AE Sensor Not Recognized After Multiple Reconnects cycles and Orange LED Remains Active | Intermittent Connectivity | Reconnection of sensor not stable | Christian | yes |  |  |  | Christian Kurz | yes (with device logs)
Y3839-661 | SensorConnected property does not update in Upgrade State | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | FW bug |  |  | Michael Petek | no | yes | 1 week | new FW
Y3839-662 | HyperRAM is not erased for Gen3 Sensors | Previous (Patient) Image | Missing error handling in SW/FW | Steffen | Yes | No checkerboard support for AE sensor |  |  | Xiang Ding
Y3839-663 | Power cycle leads to unstable voltage | Unassigned |  | Christian | DyingBoxes | Others currently unknown |  | Maybe this "irregular" power state lead to dying boxes? | Frank Darmer
Y3839-664 | Broken FPGA image still loaded after sensor unplugged | Previous (Patient) Image | Missing error handling in SW/FW | Steffen | Interface Update Issues | FW bug |  |  | Michael Petek | no | yes | 1 week | new FW
Y3839-673 | Updater Tool: No update possible if no USB drivers are installed | Interface Update Issues | Driver conflicts after IOSS/CDR reinstall for 3rd party | Michael | yes | IOSS not installed |  |  | Team_DX5
Y3839-689 | Updater Tool: Downgrade cannot be re-triggered after the FPGA flash step fails | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | Firmware Upgrader bug |  |  | Team_DX5
Y3839-690 | Error flag destroys USB3 channel to host computer | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes |  |  |  | Michael Petek | no | yes | 1 week | new FW
Y3839-691 | Sensor detection not stable | Intermittent Connectivity | Reconnection of sensor not stable | Christian | yes |  |  | reassigned to Xiang, initial sensor detection is handled by FPGA | Xiang Ding
Y3839-695 | Unexpected Orange Residue Present on Electronic Chips of USB3 Interfaces After Manufacturing | No Power / "Dying Boxes" (USB module) | Others currently unknown | Guido | no | Y3839-695 has no impact to the electronic |  |  | Frank Darmer
Y3839-696 | Preview image is not always available within 7seconds as required | 3rd Party Slowness (NAM) | Others currently unknown | Erik | Yes | general slowness |  |  | Team_DX5 | no | yes | 0 weeks (done) | new IOSS | 5
Y3839-698 | On plug-in of incompatible sensor, Sensor LED lights up green before turning red | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync) | Erik | No | FW Led state bug |  | reassigned to Xiang, sensor LED is controlled by FPGA | Xiang Ding
Y3839-699 | Odontogram in indication dialog is active only if it's not predefined in viewset | Unassigned |  | Steffen | not assigned | documentation mismatch |  |  | Christian Beckhaus
Y3839-702 | Indication region in Sidexis is different to the one in indication dialog | Interface Update Issues | Missing error handling in SW/FW | Michael | no | cannot be categorized, Docu bug / IOSS bug |  |  | Christian Beckhaus
Y3839-705 | IOSS Installer: Error Text Stays German Even if other languages is selected | Interface Update Issues | Missing error handling in SW/FW | Michael | no | cannot be categorized, IOSS Installer bug |  |  | Team_DX5
Y3839-706 | Documentation: rotation buttons are usable | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync) | Erik | No | Software / UI Bug |  |  | Christian Beckhaus
Y3839-709 | Endo horizontal and endo vertical templates do not contain predefined anatomic regions | Unassigned |  | Christian | not assigned | documentation mismatch |  |  | Christian Beckhaus
Y3839-710 | Remote Box is not usable after reinstall of IOSS/CDR Drivers | Interface Update Issues | Driver conflicts after IOSS/CDR reinstall for 3rd party | Michael | no | cannot be categorized, IOSS driver reinstallation |  |  | Erik Roecher | done | yes | 0.5 weeks | new IOSS | 2
Y3839-711 | AE Updater tool blocks update due to false sensor connection detection | Interface Update Issues | Missing error handling in SW/FW | Michael | yes | Firmware Upgrader bug |  |  | ?
Y3839-712 | Documentation: Recovery references multiple unfinished acquisition sessions although only one session can exist | Previous (Patient) Image | Workflow robustness | Steffen | not assigned | documentation |  |  | Christian Beckhaus
Y3839-718 | Corrupted Calibration file | Intermittent Connectivity | Calibration file not transferred correctly | Christian | yes |  |  |  | Xiang Ding
Y3839-722 | CDR: Sensor remains armed (event detection not disabled) if integrator transitions from State 6 directly to State 3 (Idl | White Images | Sensor permanently armed (scattering X-Ray, background accumulation) | Xiang | Yes |  |  |  | Michael Petek | no | yes, but discussed with SW (only for exposure readiness) | 1 week | new FW, potentially new IOSS
Y3839-732 | Sensor stays unregistered in Sidexis after 401 on device re-registration (expired refresh token, no retry) | Intermittent Connectivity | Missing error handling in SW/FW | Christian | yes | IOSS bug? |  |  | Steffen Wenz | no | yes | 0.3 weeks | new IOSS | 3
Y3839-733 | CDR AutoAcquire setting has no impact in Eaglesoft | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | Yes |  |  | CDR not IOSS | ? | no | no |  | is a CDR specific issue
Y3839-734 | Synergy/Raptor USB Interface Module Self-Resets and Disconnects During Acquisition (reproducible only on the USB 2.0 box | Unassigned |  | Christian | ? | USB2 box |  |  | ?
Y3839-736 | IOSS - Eaglesoft: Bitewing Series - same exposure is loaded in two frames and stuck at taking the exposure for the third | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW | Erik | Yes |  |  |  | Albert Heller
New added tickets
Y3839-Y731 | Inconsistent Blinking of LEDs at boot | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync) |  | no | Inconstant ready-for-exposure signaling (SW vs. Interface) | FW bug | Michael Petek
Y3839-Y738 | Unplugging Sensor in State Image Ready Leads To Image Loss | Images not transferred | State-machine mismatch in SW/FW (out-of-sync) |  | no | Images not transferred | Missing error handling in SW/FW | Michael Petek
Y3839-Y739 | Interface goes into reboot loop when Gen 2 sensor is plugged | Interface Update Issues | Missing error handling in SW/FW |  | no | intermittent connectivity | others currently unknown | Xiang Ding
Y3839-Y741 | Sensor LED is green even if sensor is not responding | Inconstant ready-for-exposure signaling (SW vs. Interface) | Missing error handling in SW/FW |  | no | Inconstant ready-for-exposure signaling (SW vs. Interface) | FPGA Bug | Xiang Ding
Y3839-Y742 | Race Conditions during GPIF transactions | Images not transferred | State-machine mismatch in SW/FW (out-of-sync) |  | no | 3rd Party Slowness (NAM) | general slowness | Michael Petek
Y3839-Y743 | Interface Resets On Sensor Plug And Unplug | Intermittent Connectivity | Reconnection of sensor not stable |  | yes |  |  | Michael Petek
Y3839-Y744 | Dentrix/CDR: USB3 sporadically fails to detect sensor after repeated plug/unplug – Sensor LED stays off, device unresponsive to reconnection | Intermittent Connectivity | Reconnection of sensor not stable |  | yes |  |  | Steffen Wenz
Y3839-Y745 | FPGA dies after interface lies around with plugged sensor | Intermittent Connectivity | Missing error handling in SW/FW |  | no | Intermittent Connectivity | FPGA Bug | Xiang Ding

## Trainingsdaten
Ticket | expected symptom | expected rootcause
Y3839-231 | Intermittent Connectivity | Reconnection of sensor not stable
Y3839-234 | No Power / "Dying Boxes" (USB module) | USB3 remote hardware failure
Y3839-255 | Images not transferred | Missing error handling in SW/FW
Y3839-265 | not assigned
Y3839-282 | Images not transferred | Missing error handling in SW/FW
Y3839-304 | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue
Y3839-305 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-307 | Interface Update Issues | Missing error handling in SW/FW
Y3839-308 | not assigned
Y3839-309 | not assigned
Y3839-340 | not assigned
Y3839-373 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-382 | IOSS Bugs
Y3839-383 | Intermittent Connectivity | Loose micro-B USB interface (no cable clip used)
Y3839-384 | Images not transferred | IOSS Bug
Y3839-422 | not assigned
Y3839-425 | not assigned
Y3839-426 | not assigned
Y3839-471 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-478 | Intermittent Connectivity | IOSS Bug
Y3839-479 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-506 | Interface Update Issues | IOSS Bug
Y3839-510 | Interface Update Issues | Missing error handling in SW/FW
Y3839-511 | not assigned
Y3839-513 | Intermittent Connectivity | IOSS Bug
Y3839-515 | not assigned
Y3839-520 | Images not transferred | IOSS Bug
Y3839-521 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-522 | Intermittent Connectivity | IOSS Bug
Y3839-523 | Previous (Patient) Image | IOSS Bug
Y3839-524 | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync)
Y3839-525 | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW timing issue
Y3839-526 | --> in my opinion solved ?
Y3839-529 | Images not transferred | IOSS Bug
Y3839-530 | unassigned
Y3839-533 | Inconstant ready-for-exposure signaling (SW vs. Interface) | Missing error handling in SW/FW
Y3839-537 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-538 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-541 | Previous (Patient) Image | Others currently unknown
Y3839-545 | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync)
Y3839-546 | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync)
Y3839-547 | Images not transferred | SW timing issue
Y3839-549 | Loosening Screws | Incorrect torque used during sensor cable exchange
Y3839-555 | Images not transferred | Sensor self-triggering
Y3839-567 | Inconstant ready-for-exposure signaling (SW vs. Interface) | State-machine mismatch in SW/FW (out-of-sync)
Y3839-568 | Intermittent Connectivity | Missing Error handling in SW/FW
Y3839-570 | Images not transferred | State-machine mismatch in SW/FW (out-of-sync)
Y3839-575 | Intermittent Connectivity | Reconnection of sensor not stable
Y3839-586 | White Images | Sensor self-triggering
Y3839-593 | Images not transferred | Missing error handling in SW/FW
Y3839-594 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-596 | White Images | Sensor triggers by the influence of an electric field
Y3839-598 | Intermittent Connectivity | Driver conflicts (CDR vs IOSS)
Y3839-599 | Intermittent Connectivity | Reconnection of sensor not stable
Y3839-607 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-611 | Interface Update Issues | FW Bug
Y3839-625 | Interface Update Issues | FPGA Bug
Y3839-631 | Inconstant ready-for-exposure signaling (SW vs. Interface) | FW Bug
Y3839-639 | Inconstant ready-for-exposure signaling (SW vs. Interface) | FW Bug
Y3839-657 | Intermittent Connectivity | Reconnection of sensor not stable
Y3839-661 | Interface Update Issues | FW Bug
Y3839-662 | Previous (Patient) Image | No checkerboard support for AE sensor
Y3839-663 | DyingBoxes | Others currently unknown
Y3839-664 | Interface Update Issues | FW Bug
Y3839-673 | Interface Update Issues | IOSS not installed
Y3839-689 | Interface Update Issues | FW Bug
Y3839-690 | Intermittent Connectivity | Missing error handling in SW/FW
Y3839-691 | Intermittent Connectivity | Reconnection of sensor not stable
Y3839-695 | Y3839-695 has no impact to the electronic
Y3839-696 | 3rd Party Slowness (NAM) | general slowness
Y3839-698 | Inconstant ready-for-exposure signaling (SW vs. Interface) | FW Bug
Y3839-699 | not assigned
Y3839-702 | not assigned
Y3839-705 | not assigned
Y3839-706 | Inconstant ready-for-exposure signaling (SW vs. Interface) | SW Bug
Y3839-709 | not assigned
Y3839-710 | not assigned
Y3839-711 | Interface Update Issues | FW Bug
Y3839-712 | not assigned
Y3839-718 | Intermittent Connectivity | Calibration file not transferred correctly
Y3839-722 | White Images | Sensor permanently armed (scattering X-Ray, background accumulation)
Y3839-732 | Intermittent Connectivity | IOSS Bug
Y3839-733 | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW
Y3839-734 | not assigned
Y3839-736 | Inconstant ready-for-exposure signaling (SW vs. Interface) | Workflow sequence not correctly used by 3rd party SW
Y3839-Y731 | Inconstant ready-for-exposure signaling (SW vs. Interface) | FW bug
Y3839-Y738 | Images not transferred | Missing error handling in SW/FW
Y3839-Y739 | intermittent connectivity | others currently unknown
Y3839-Y741 | Inconstant ready-for-exposure signaling (SW vs. Interface) | FPGA Bug
Y3839-Y742 | 3rd Party Slowness (NAM) | general slowness
Y3839-Y744 | Intermittent Connectivity | Reconnection of sensor not stable
Y3839-Y745 | Intermittent Connectivity | FPGA Bug

## Release-Scope
How can the Release Scope look like
topic | components
stabilize sensor detection (fast connection, slow connection gen 1) | FPGA - FW
stabilize calibration data transfer | FPGA - FW -SW
utilize sensor keep alive for sensor connection check | FPGA - FW
stabilize status information (e.g. polling, periodical message from remote) | FW - SW
disarm sensor after each X-Ray image | FPGA - FW
improve/modify IOSS-API, that only a correct utilization from 3rd party is possible (enable event detection -> breathing of interface -> X-Ray image is possible) | SW
check power supply | HW
check EMI on image transfer | HW

## R&D Improvements
Twain IOSS support without SIdexis
Datadog connectin in IOSS
Only provide Updates via IOSS (OTA)
Prevent CDR customers to use new Firmwares (to allow interface changes)
(hidden) option to save raw images for image quality studies or bugfixing
[Y3839-737] Check for CDR Autodetect Monitor - Jira Dentsply Sirona
Log Einträge der USB Box auf dem PC verfügbar machen
IOSS Rollout via MS Store
Note on A6: Beckhaus, Christian: in IOSS